## filter_wheel_mounting
Mounted Science Filter | Combination Filter | Passband (90%) (nm) | Central wavelength (nm) | Optimal Tilt (deg) | Available ring (deg) (Red color indicates the ring is chosen based on availability of rings instead of spectroscopic testing) | Remarks | Slot | Status | Remarks | Visual Inspection | Available ring tilt (deg) | Quantity (FM+QM) | Filters | Rings left
NB1_1 | BB1 | 200–235.8 | 214 |  | 6 | No major variation of transmission with tilt. | FW2_5 | Mounted. All torqued. | Bolt 6 changed to 14mm. | Bolt face (sun side): DHA face: | 0 | 4 | NB8, NB8
NB2A_7 | BP2 | 275.94–277.46 | 276.6 | 6 | 6 |  | FW1_6 | Mounted. All torqued at 12Ncm with glue. | NB2A_10 cracked at 20Ncm | Bolt face (sun side): Scattered lint. DHA face: Imprints, Small dot like reddish appearance. | 1 | 4 |  | 4
NB3_3 | BP2 | 277.595–281.605 | 279.6 | 5 | 5 |  | FW1_5 | Mounted. All torqued. |  | Bolt face (sun side): Scattered dust. DHA face: Imprint, wipe mark, little lint. | 2 | 8 |  | 8
NB4_2 | BP2 | 278.62–281.98 | 280.3 | 5 | 5 |  | FW1_4 (HOME) | Mounted. All torqued at 20Ncm. Slight crack like appearance on the side. Bolts potted in a semicircle around bolt head. | NB4_1 cracked | Bolt face (sun side): Crack along filter edge. Scattered dust and lint. DHA face: Little scattered dust. | 4 | 6 | NB6, BB1, BB1, BB3, NB7 | 1
NB5_4 | BP2 | 282.25–284.15 | 283.2 | 5 | 5 |  | FW1_7 | Mounted. All torqued. |  | Bolt face (sun side): Scattered dust. DHA face: Little imprint, Little scattered dust. | 5 | 6 | NB3, NB4, NB5, BP3, BP4 | 1
NB6_3 | BP3 | 297.45–302.55 | 300 | 3 | 4 | 2 or 4 may have to be used. | FW2_1 | Mounted. All torqued. |  | Bolt face (sun side): DHA face: | 6 | 6 | NB2A, NB1, BP2, BB2 | 2
NB7_1 | BP3 | 384.35–391.65 | 388 | 3 | 4 | 4 may be used instead of 3 | FW2_2 | Mounted. All torqued.
NB8_1 | NB8 | 395.801–397.9 | 396.85 | 0 | 0 |  | FW2_3 | Mounted. All torqued. | NB8_3 cracked
BB1_5 | BB1 | 200–242 |  |  | 4 |  | FW1_2 | Mounted. All torqued at 12Ncm with glue under bolt head. | BB1_1 cracked, BB1_2 cracked. Cracking at 20Ncm. | Bolt face (sun side): Scratch marks. Rough appearance on coated surface. DHA face: Wipe marks. Rough appearance on coated surface.
BB2_3 | BP4 | 242–300 |  |  | 6 |  | FW2_6 | Mounted. All torqued.
BB3_3 | BP4 | 320–360 |  |  | 4 |  | FW2_7 | Mounted. All torqued.
Combination Filters.
BP2_8 |  |  |  |  | 6 |  | FW2_8 (HOME) | Mounted. 1-3, 5-6 torqued. Point 4 thread damaged- No bolt at 4. Point 2 bolt very tight in helicoil, glued bolt head to lock ring. | BP2_6 cracked. Point 4 was slipping. Was anticipated to be a tool problem.
BP3_2 |  |  |  |  | 5 |  | FW1_8 | Mounted. All torqued. |  | Bolt face (sun side): Little scattered lint. DHA face: Little scatterd dust.
BP4_4 |  |  |  |  | 4 |  | FW1_3 | Mounted. All torqued. |  | Bolt face (sun side): Scattered dust and lint. DHA face: Scattered dust. Little mark.
BB1_3 |  |  |  |  | 4 |  | FW2_4 | Mounted. All torqued. Small crack on sun side. | BB1_1 cracked, BB1_2 cracked. Cracking at 20Ncm while placing on FW1.
NB8_2 |  |  |  | 0 | 0 |  | FW1_1 | Mounted. All torqued at 20Ncm. Bolts potted in a semicircle around bolt head. |  | Bolt face (sun side): Slight lint, dust, wipe mark. DHA face: Wipe marks.
Sorted as per FW slot
FW1
NB8_2 | 0 | 0 | 0 | 0 | 0 | 0 | FW1_1 | Mounted. All torqued at 20Ncm. Bolts potted in a semicircle around bolt head. | 0
BB1_5 | BB1 | 200–242 | 0 |  | 4 | 0 | FW1_2 | Mounted. All torqued at 12Ncm with glue under bolt head. | BB1_1 cracked, BB1_2 cracked. Cracking at 20Ncm.
BP4_4 | 0 | 0 | 0 |  | 4 | 0 | FW1_3 | Mounted. All torqued. | 0
NB4_2 | BP2 | 278.62–281.98 | 280.3 | 5 | 5 | 0 | FW1_4 (HOME) | Mounted. All torqued at 20Ncm. Slight crack like appearance on the side. Bolts potted in a semicircle around bolt head. | NB4_1 cracked
NB3_3 | BP2 | 277.595–281.605 | 279.6 | 5 | 5 | 0 | FW1_5 | Mounted. All torqued. | 0
NB2A_7 | BP2 | 275.94–277.46 | 276.6 | 6 | 6 | 0 | FW1_6 | Mounted. All torqued at 12Ncm with glue. | NB2A_10 cracked at 20Ncm
NB5_4 | BP2 | 282.25–284.15 | 283.2 | 5 | 5 | 0 | FW1_7 | Mounted. All torqued. | 0
BP3_2 | 0 | 0 | 0 |  | 5 | 0 | FW1_8 | Mounted. All torqued. | 0
FW2
NB6_3 | BP3 | 297.45–302.55 | 300 | 3 | 4 | 2 or 4 may have to be used. | FW2_1 | Mounted. All torqued. | 0
NB7_1 | BP3 | 384.35–391.65 | 388 | 3 | 4 | 4 may be used instead of 3 | FW2_2 | Mounted. All torqued. | 0
NB8_1 | NB8 | 395.801–397.9 | 396.85 | 0 | 0 | 0 | FW2_3 | Mounted. All torqued. | NB8_3 cracked
BB1_3 | 0 | 0 | 0 |  | 4 | 0 | FW2_4 | Mounted. All torqued. Small crack on sun side. | BB1_1 cracked, BB1_2 cracked. Cracking at 20Ncm while placing on FW1.
NB1_1 | BB1 | 200–235.8 | 214 |  | 6 | No major variation of transmission with tilt. | FW2_5 | Mounted. All torqued. | Bolt 6 changed to 14mm.
BB2_3 | BP4 | 242–300 | 0 |  | 6 | 0 | FW2_6 | Mounted. All torqued. | 0
BB3_3 | BP4 | 320–360 | 0 |  | 4 | 0 | FW2_7 | Mounted. All torqued. | 0
BP2_8 | 0 | 0 | 0 |  | 6 | 0 | FW2_8 (HOME) | Mounted. 1-3, 5-6 torqued. Point 4 thread damaged- No bolt at 4. Point 2 bolt very tight in helicoil, glued bolt head to lock ring. | BP2_6 cracked. Point 4 was slipping. Was anticipated to be a tool problem.
1
2
4
3
5
6
0
FW ring screw positions

## FW_slot_diagram
FW1 as seen from Thermal filter side
Slot Number as marked on FW
Tilt Angle
Filter Name-Number
Filter Number in
Program sequence
4
5°
NB4-2
00
5
5°
NB3-3
01
6
6°
NB2A-7
02
7
5°
NB5-4
03
8
5°
BP3-2
04
2
4°
BB1-5
06
3
4°
BP4-4
07
1
0°
NB8-2
05
FW1
Home
FW2 as seen from DHA side
Home
Slot Number as marked on FW
Tilt Angle
Filter Name-Number
Filter Number in
Program sequence
8
6°
BP2-8
00
1
4°
NB6-3
01
2
4°
NB7-1
02
3
0°
NB8-1
03
4
4°
BB1-3
04
6
6°
BB2-3
06
7
4°
BB3-3
07
5
6°
NB1-1
05
FW2

## Test_matrix
Filter Name | Available Sl. No. | Shelf | Date of Spatial Test | Date of OOB Test | Date of Tilt Test | Remarks | Legend: | OK | FM | Not OK/ Not Tested
NB1 | 2 QM | Top | 2022-10-10 | 2022-10-10 | 2022-10-26 | Trans. lower than vendor but consistent in all spatial direction
NB1 | 1 FM |  | 2023-01-31 |  |  | OOB of FM filters are not taken due to contamination risk, as instructed previously. OOB of all QM are to be recorded.
NB2A | 7 QM | Top | 2022-09-08 | 2022-09-08 | 2022-10-27
NB2A | 8 | Top | 2022-09-21 | 2022-09-21
NB2A | 10 FM | Top | 2022-09-21 | 2022-09-21
NB3 | 1 | Mid | 2022-09-22 | 2022-09-22
NB3 | 2 QM | Mid | 2022-04-06 | 2022-04-06 | 2022-10-17
NB3 | 3 FM | Mid | 2022-04-07
NB4 | 1 FM | Top | 2022-04-07
NB4 | 2 QM | Top | 2022-04-06 | 2022-04-06 | 2022-10-17
NB4 | 3 | Top | 2022-09-23 | 2022-09-23
NB4 | 4 | Top | 2022-09-23 | 2022-09-23
NB4 | 5 | Top | 2022-09-23 | 2022-09-23
NB5 | 1 | Top
NB5 | 3 QM | Top | 2022-09-09 | 2022-09-09 | 2022-10-27
NB5 | 4 FM | Top | 2023-01-31/2023-02-02
NB6 | 1 QM | Mid |  | 2023-11-11 | 2022-10-28
NB6 | 2 | Mid
NB6 | 3 FM | Mid | 2022-04-07
NB6 | 4 | Mid
NB7 | 1 FM | Mid | 2022-04-07
NB7 | 2 QM | Mid |  | 2023-11-11 | 2022-10-28
NB7 | 3 | Mid
NB7 | 4 | Mid
NB8 | 1 QM | Top | 2022-04-25 | 2022-04-25 |  | Different spatial transmission. Not matching with Vendor
NB8 | 2 FM | Top | 2023-01-31/ 2023-02-02 |  |  | Different spatial transmission. Not matching with Vendor
NB8 | 3 FM | Top | 2023-01-31/ 2023-02-02 |  |  | Different spatial transmission. Not matching with Vendor
NB8 | 4 QM | Top
NB8 | 5 | Top
BB1 | 1 FM | Bottom | 2023-01-30 | 2023-01-30 |  | OOB not less than 0.01% due to second and higher peaks.
BB1 | 2 QM | Bottom | 2022-10-10 | 2022-10-10 |  | Trans. lower than vendor but consistent in all spatial direction
BB1 | 3 FM | Bottom | 2023-01-30 | 2023-01-30 |  | OOB not less than 0.01% due to second and higher peaks.
BB1 | 4 | Bottom
BB1 | 5 QM | Bottom | 2022-10-07 | 2022-10-07 |  | Trans. lower than vendor but consistent in all spatial direction
BB2 | 1 | Bottom
BB2 | 2 | Bottom
BB2 | 3 FM | Bottom | 2023-01-28 | 2023-01-28 |  | OOB not less than 0.01% due to second and higher peaks.
BB2 | 4 QM | Bottom | 2022-10-06
BB2 | 5 | Bottom
BB3 | 1 | Bottom
BB3 | 2 | Bottom
BB3 | 3 FM | Bottom | 2023-01-28 | 2023-01-28 |  | OOB not less than 0.01% due to second and higher peaks.
BB3 | 4 QM | Bottom | 2022-09-13 | 2022-09-13 |  | BUMP in OOB may be due to spectrograph
BB3 | 5 | Bottom
BP2, Now BP1 | 6 FM | Top | 2023-01-29 | 2023-01-29 |  | Our Trans. Higher than Vendor and OOB not less than 0.01% due to second order.
BP2, Now BP1 | 7 | Top
BP2, Now BP1 | 8 QM | Top | 2022-04-26 | 2022-04-26
BP2, Now BP1 | 9 | Top
BP2, Now BP1 | 10 | Top
BP3, Now BP2 | 1 QM | Bottom | 2022-09-09 | 2022-09-09
BP3, Now BP2 | 2 FM | Bottom | 2023-01-29 | 2023-01-29 |  | OOB not less than 0.01% due to second order.
BP3, Now BP2 | 7 | Bottom
BP3, Now BP2 | 9 | Bottom
BP3, Now BP2 | 10 | Bottom
BP4, Now BP3 | 1 | Bottom
BP4, Now BP3 | 2 QM | Bottom | 2022-04-26 | 2022-04-26
BP4, Now BP3 | 3 | Bottom
BP4, Now BP3 | 4 FM | Bottom | 2023-01-29 |  |  | OOB not taken due to long exp time requirements. See readme.
BP4, Now BP3 | 5 | Bottom

## OLD_test_matrix
FILTER DATA LOG: SERIAL NUMBERS UPDATED FOR QM AND FM BASED ON AVAILABILITY.
Filter Name | Available Sl. No. | Shelf | QM sl.no. | Date of Spectral Transmission | Spectral Transmission remarks | Date of QM Spatial and OOB Test | FM sl.no. | Date of Spectral Transmission | Spectral Transmission Remarks | Date of FM Spatial Test
NB1 | 1,2,3,4,5 | Top | 2 | 2022-02-21 |  |  | 1 | 2022-03-18 | Data discrepancy seen on 2022-03-16
NB2A | 7,8,10 | Top | 7 | 2022-02-21 |  | 2022-09-08 | 10
NB3 | 1,2,3 | Mid | 2 | 2022-03-09 |  | 2022-04-06 | 3 | 2022-03-16 | Photo taken (161409) | 2022-04-07
NB4 | 1,2,3,4,5 | Top | 2 | 2022-03-09 |  | 2022-04-06 | 1 | 2022-03-16 | Photo taken (155637) | 2022-04-07
NB5 | 1,3,4 | Top | 3 | 2022-02-21 |  | 2022-09-09 | 4
NB6 | 1,2,3,4 | Mid | 1 | 2022-02-21 |  |  | 3 | 2022-03-16 |  | 2022-04-07
NB7 | 1,2,3,4 | Mid | 2 | 2022-02-21 |  |  | 1 | 2022-03-16 |  | 2022-04-07
NB8 | 1,2,3,4,5 | Top | 4,1 | 09-03-2022 (Filter #1) | (Data taken for #1) Out of Band also taken. | 2022-04-25 (Filter #1) To be repeated. | 2, 3 | 16-03-2022 (filter #3) | Photo of #3 taken (163740 and 163811)
BB1 | 1,2,3,4,5 | Bottom | 2,5 | 2022-03-09 |  |  | 1,3 | 2022-03-17
BB2 | 1,2,3,4,5 | Bottom | 4 | 2022-02-24 |  |  | 3 | 2022-03-18
BB3 | 1,2,3,4,5 | Bottom | 4 | 2022-02-24 |  | 04/26/2022 09/13/2022 To be repeated (OOB bump) | 3 | 2022-03-17
BP2, Now BP1 | 6,7,8,9,10 | Top | 8 |  |  | 05/25/2022 To be repeated? (Higher transmission) | 6
BP3, Now BP2 | 1, 2, 9, 7, 10 | Bottom | 1 |  |  | 2022-09-09 | 2 | 2022-03-09 | Done
BP4, Now BP3 | 1-5 | Bottom | 2 |  |  | 04/26/2022 To be repeated (OOB Bump) | 4
BP1, Now Cancelled |  |  |  |  | Full series N/A
NB2 | 2,4 | Mid |  |  |  | 2022-09-08 (Nidhi)
Filter choice as given by Soumya-
[6 empty] | Filter Name | QM sl.no. | FM sl.no.
[6 empty] | NB1 | 2 | 1
[6 empty] | NB2A |  | 7
NB2
[6 empty] | NB3 | 2 | 3
[6 empty] | NB4 | 2 | 1
[6 empty] | NB5 |  | 3
[6 empty] | NB6 | 1 | 3
[6 empty] | NB7 | 2 | 1
[6 empty] | NB8 | 4,1 | 2, 3
[6 empty] | BB1 | 2,5 | 1,3
[6 empty] | BB2 | 4 | 3
[6 empty] | BB3 | 4 | 3
BP1, Now Cancelled
[6 empty] | BP2, Now BP1 | 8 | 6
[6 empty] | BP3, Now BP2 | 1 | 3
[6 empty] | BP4 (340nm/100nm), Now BP3 | 2 | 4
[6 empty] | Unavailable filters marked in red- | NOTE: THE QM FITLERS HAVE BEEN MARKED AS FM AS SUGGESTED BY SOUMYA
[6 empty] | Filter Name | Sl. No. available | QM | FM
[6 empty] | NB2A | 7,8,10 | 7 | 6
[6 empty] | NB5 | 1,3,4 | 3 | 2
[6 empty] | BP3 | 1, 2, 9, 7, 10 | 1 | 3